## Data1
[8 empty] | PAGE 1 OF |  | 2 PAGES
CALIBRATION DATA
Calibrated By :
Calibration Date | : | 2016-01-27 | [6 empty] | CERTIFICATE NUMBER : | PSA/BMCL/16/018-NA
X
Issued By | : | WHIZZ SYSTEMS MALAYSIA SDN. BHD. | [6 empty] | Cal. Sticker No. : | 18
[12 empty] | J | Ramli B. Jiman
Issued To | : | [7 empty] | Job No. : | BMCL/15/022-4 |  | H | Nor Sham Bt. Hasan
[9 empty] | Date Received : |  |  | M | Liyana Bt. Mahmud
[12 empty] | A | Roslina Bt. Abdul Aziz
[12 empty] | Y | Rusnani Bt. Yahya
[12 empty] | G | Gopi A/L Muthusamy
X
Instrument Name | : | Check Master | [6 empty] | Model No. : |  |  | Ö
Manufacturer | : | [7 empty] | Serial No. :
Instrument Condition When Received : |  |  | Tick | [5 empty] | Start | End
1. Good physical condition. |  |  |  |  | Time |  |  | : |  |  | am / pm
2. (*) Denote out of test limits.
3. …………………………….. |  |  |  |  | Environment Conditions :
[9 empty] | Min | Max
Instrument Condition When Returned : |  |  |  |  | Ambient Temperature |  |  | : |  |  | °C
1. Calibrated and tested serviceable.
2. No Adjustment done. |  |  |  |  | Tolerence |  |  | : | 0 | °C
3. (*) Denote out of test limits.
4. Performance Limitation. | [8 empty] | Min | Max
5. Adjustment was carried out. |  |  |  |  | Relative Humidity |  |  | : |  |  | %
6. Calibrated at Customer Specific Range/Method.
7. Requested Due Date : |  | - Year |  |  | Tolerence |  |  | : | 0 | %
8. ........................................
Calibration Method :
This instrument was calibrated using the Calibration Procedure No. : | [5 empty] | BMCL/CP/02 Rev.0
Calibration Standard Used :
Instrument Type : | Serial No. : |  | Cal. Due Date : |  |  | Cal. Cert. No. : |  |  | Traceability :
HEISE ST-2H (HQS-2)
CALIBRATION DATA / CERT STATUS CHECK
|  | Checked By |  | Date |  |  | Remark
CAL DATA
Note on N9: PSA/BMCL/YR/xxx
Note on N11: xxxx
Note on N14: BMCL/YR/xxx

## Data2
[11 empty] | PAGE 2 OF | 2 PAGES
CALIBRATION DATA
Calibration Date : |  | 2016-01-27 | [8 empty] | Certificate No. : | PSA/BMCL/16/018-NA
Instrument : |  | Check Master | [8 empty] | Serial No. :
0
Function / Range |  |  | Parameter Test |  |  | Tolerance |  |  | Test Result (Pass / Fail)
1. Pressure Leak |  |  | 400 mmHg |  |  | < 2 mmHg/min
Function / Range |  |  | MS Setting |  | Test Limits |  |  |  | Measured Value
[5 empty] | Minimum |  | Maximum |  | Test 1 | Test 2 | Test 3 | Avg | Unit | Diff (Test 1- Test 2) | Diff Sqr | Certificate (1 psi = 51.7149 mmHg) | Spec (0.05% x Avg) / k
[18 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
2. Pressure Static Test
Increase Value
| 0 | mmHg | 0 | mmHg | -1 | mmHg | 1 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 0 | 1 | -1 | 1 | #DIV/0!
| 50 | mmHg | 50 | mmHg | 48.75 | mmHg | 51.25 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 50 | 1.25 | 48.75 | 51.25 | #DIV/0!
| 100 | mmHg | 100 | mmHg | 98.5 | mmHg | 101.5 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 100 | 1.5 | 98.5 | 101.5 | #DIV/0!
| 150 | mmHg | 150 | mmHg | 148.25 | mmHg | 151.75 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 150 | 1.75 | 148.25 | 151.75 | #DIV/0!
| 200 | mmHg | 200 | mmHg | 198 | mmHg | 202 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 200 | 2 | 198 | 202 | #DIV/0!
| 250 | mmHg | 250 | mmHg | 247.75 | mmHg | 252.25 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 250 | 2.25 | 247.75 | 252.25 | #DIV/0!
| 300 | mmHg | 300 | mmHg | 297.5 | mmHg | 302.5 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 300 | 2.5 | 297.5 | 302.5 | #DIV/0!
| 350 | mmHg | 350 | mmHg | 342 | mmHg | 358 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 2 | 1 | 1 | 350 | 8 | 342 | 358 | #DIV/0!
| 400 | mmHg | 400 | mmHg | 391 | mmHg | 409 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 2 | 1 | 1 | 400 | 9 | 391 | 409 | #DIV/0!
Decrease Value
| 400 | mmHg | 400 | mmHg | 391 | mmHg | 409 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 2 | 1 | 1 | 400 | 9 | 391 | 409 | #DIV/0!
| 350 | mmHg | 350 | mmHg | 342 | mmHg | 358 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 2 | 1 | 1 | 350 | 8 | 342 | 358 | #DIV/0!
| 300 | mmHg | 300 | mmHg | 297.5 | mmHg | 302.5 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 300 | 2.5 | 297.5 | 302.5 | #DIV/0!
| 250 | mmHg | 250 | mmHg | 247.75 | mmHg | 252.25 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 250 | 2.25 | 247.75 | 252.25 | #DIV/0!
| 200 | mmHg | 200 | mmHg | 198 | mmHg | 202 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 200 | 2 | 198 | 202 | #DIV/0!
| 150 | mmHg | 150 | mmHg | 148.25 | mmHg | 151.75 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 150 | 1.75 | 148.25 | 151.75 | #DIV/0!
| 100 | mmHg | 100 | mmHg | 98.5 | mmHg | 101.5 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 100 | 1.5 | 98.5 | 101.5 | #DIV/0!
| 50 | mmHg | 50 | mmHg | 48.75 | mmHg | 51.25 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 50 | 1.25 | 48.75 | 51.25 | #DIV/0!
| 0 | mmHg | 0 | mmHg | -1 | mmHg | 1 | mmHg |  |  |  | #DIV/0! | mmHg | 0 | 0 | 0.10343 | #DIV/0! | 0.5 | 1 | 1 | 0 | 1 | -1 | 1 | #DIV/0!
[14 empty] | Sum Dif Sqr | 0
[14 empty] | N | 18
Function / Range |  |  | Parameter Test |  |  | Tolerance |  |  | Test Result (Pass / Fail)
3. Simulation Test |  |  | 120/80 (90) |  |  | ± 5 mmHg
Calibrated by : | [7 empty] | Checked by :
Date : |  | 2016-01-27 | [5 empty] | Date :
Note on U18: k factor from cert = 2.00 DOF or N = 18 due to no of reading