## Data1
[8 empty] | PAGE 1 OF |  | 3 PAGES
CALIBRATION DATA
Calibrated By :
Calibration Date | : | 2016-01-07 | [6 empty] | CERTIFICATE NUMBER : | PSA/BMCL/16/015
X
Issued By | : | CENTURY CHEMICAL WORKS SENDIRIAN BERHAD | [6 empty] | Cal. Sticker No. : | 15
[12 empty] | J | Ramli B. Jiman
Issued To | : | [7 empty] | Job No. : | BMCL/16/002-1 |  | H | Nor Sham Bt. Hasan
[9 empty] | Date Received : |  |  | M | Liyana Bt. Mahmud
[12 empty] | A | Roslina Bt. Abdul Aziz
[12 empty] | Y | Rusnani Bt. Yahya
[12 empty] | G | Gopi A/L Muthusamy
X
Instrument Name | : | Digital Multimeter | [6 empty] | Model No. : | 105 |  | Ö
Manufacturer | : | Appa | [6 empty] | Serial No. : | 2016-ID-3
Instrument Condition When Received : |  |  | Tick | [5 empty] | Start | End
1. Good physical condition. |  |  |  |  | Time |  |  | : |  |  | am / pm
2. (*) Denote out of test limits.
3. …………………………….. |  |  |  |  | Environment Conditions :
[9 empty] | Min | Max
Instrument Condition When Returned : |  |  |  |  | Ambient Temperature |  |  | : |  |  | °C
1. Calibrated and tested serviceable.
2. No Adjustment done. |  |  |  |  | Tolerence |  |  | : | 0 | °C
3. (*) Denote out of test limits.
4. Performance Limitation. | [8 empty] | Min | Max
5. Adjustment was carried out. |  |  |  |  | Relative Humidity |  |  | : |  |  | %
6. Calibrated at Customer Specific Range/Method.
7. Requested Due Date : |  | - Year |  |  | Tolerence |  |  | : | 0 | %
8. ........................................
Calibration Method :
This instrument was calibrated using the Calibration Procedure No. : | [5 empty] | BMCL/CP/01 Rev.0
Calibration Standard Used :
Instrument Type : | Serial No. : |  | Cal. Due Date : |  |  | Cal. Cert. No. : |  |  | Traceability :
TDS 3012C | C015061 |  | 2015-04-23 |  |  | SSTS/R/2013 D-2245 |  |  | NML M'sia
A-M SYS 3000 | 57929 |  | 2015-04-09 |  |  | SSTS/R/2013 D-994 |  |  | NML M'sia
KEITHLEY 2000 | 1287660 |  | 2015-11-07 |  |  | SSTS/R/2013K-426 |  |  | NML M'sia
AGILENT E3631A | MY50240057 |  | 2015-04-04 |  |  | SSTS/R/2013 D-473 |  |  | NML M'sia
CALIBRATION DATA / CERT STATUS CHECK
|  | Checked By |  | Date |  |  | Remark
CAL DATA
Note on N9: PSA/BMCL/YR/xxx
Note on N11: xxxx
Note on N14: BMCL/YR/xxx

## Data2
[11 empty] | PAGE 2 OF | 3 PAGES
CALIBRATION DATA
Calibration Date : |  | 2016-01-07 | [8 empty] | Certificate No. : | PSA/BMCL/16/015
Instrument : |  | Digital Multimeter | [8 empty] | Serial No. : | 2016-ID-3
105
Function / Range |  |  | UUT |  | Test Limits |  |  |  | Measured Value
|  |  | Setting |  | Minimum |  | Maximum |  | Test 1 | Test 2 | Test 3 | Avg | Unit
1. ECG Amplitude Accuracy Test | [14 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
| 30 BPM |  | 0.5 | Vp | - |  | - |  |  |  |  | #DIV/0! | V |  | 2 |  |  | 0.5 | 0.01 | 0.49 | 0.51 | #DIV/0!
|  |  | 1 | Vp | - |  | - |  |  |  |  | #DIV/0! | V |  | 2 |  |  | 1 | 0.02 | 0.98 | 1.02 | #DIV/0!
|  |  | 1.5 | Vp | - |  | - |  |  |  |  | #DIV/0! | V |  | 2 |  |  | 1.5 | 0.03 | 1.47 | 1.53 | #DIV/0!
|  |  | 2 | Vp | - |  | - |  |  |  |  | #DIV/0! | V |  | 2 |  |  | 2 | 0.04 | 1.96 | 2.04 | #DIV/0!
2. Respiration Baseline Impedance Test | [14 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
| RL - RA |  | 525 | W | 498.75 | W | 551.25 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 525 | 26.25 | 498.75 | 551.25 | #DIV/0!
| RL - RA |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| RL - RA |  | 1500 | W | 1425 | W | 1575 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1500 | 75 | 1425 | 1575 | #DIV/0!
| RL - RA |  | 2000 | W | 1900 | W | 2100 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 2000 | 100 | 1900 | 2100 | #DIV/0!
| RL - LA |  | 525 | W | 498.75 | W | 551.25 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 525 | 26.25 | 498.75 | 551.25 | #DIV/0!
| RL - LA |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| RL - LA |  | 1500 | W | 1425 | W | 1575 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1500 | 75 | 1425 | 1575 | #DIV/0!
| RL - LA |  | 2000 | W | 1900 | W | 2100 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 2000 | 100 | 1900 | 2100 | #DIV/0!
| RL - LL |  | 525 | W | 498.75 | W | 551.25 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 525 | 26.25 | 498.75 | 551.25 | #DIV/0!
| RL - LL |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| RL - LL |  | 1500 | W | 1425 | W | 1575 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1500 | 75 | 1425 | 1575 | #DIV/0!
| RL - LL |  | 2000 | W | 1900 | W | 2100 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 2000 | 100 | 1900 | 2100 | #DIV/0!
3. Respiration V Leads Resistance Test | [14 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
| V1 - V2 |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| V1 - V3 |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| V1 - V4 |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| V1 - V5 |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
| V1 - V6 |  | 1000 | W | 950 | W | 1050 | W |  |  |  | #DIV/0! | W | #DIV/0! | 5 |  | 1 | 1000 | 50 | 950 | 1050 | #DIV/0!
Calibrated by : | [7 empty] | Checked by :
Date : |  | 2016-01-07 | [5 empty] | Date :

## Data3
[13 empty] | PAGE 3 OF | 3 PAGES
CALIBRATION DATA
Calibration Date : |  | 2016-01-07 | [10 empty] | Certificate No. : | PSA/BMCL/16/015
Instrument : |  | Digital Multimeter | [10 empty] | Serial No. : | 2016-ID-3
105
Function / Range |  |  | UUT |  | Equivalent |  | Limits |  |  |  | Measured Value
|  |  | Setting |  | Value |  | Lower |  | Upper |  | Test 1 | Test 2 | Test 3 | Avg | Unit
4. Cardiac Output Test | [16 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
| Blood Temperature |  | 36 | oC | 14.524 | kΩ | 14.514 | kΩ | 14.534 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 14.524 | 0.01 | 14.514 | 14.534 | #DIV/0!
|  |  | 37 | oC | 14.004 | kΩ | 14.002 | kΩ | 14.006 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 14.004 | 0.002 | 14.002 | 14.006 | #DIV/0!
5. Temperature Accuracy Test | [16 empty] | Spec | Digit | Res | Nom | Tol | Min | Max | Logic
| Pin 7 - Pin 2 |  | 30 | oC | 1.815 | kΩ | 1.796 | kΩ | 1.834 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 1.815 | 0.019 | 1.796 | 1.834 | #DIV/0!
| (Temp 400) |  | 35 | oC | 1.471 | kΩ | 1.456 | kΩ | 1.486 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 1.471 | 0.015 | 1.456 | 1.486 | #DIV/0!
|  |  | 37 | oC | 1.355 | kΩ | 1.341 | kΩ | 1.369 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 1.355 | 0.014 | 1.341 | 1.369 | #DIV/0!
|  |  | 40 | oC | 1.2 | kΩ | 1.188 | kΩ | 1.212 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 1.2 | 0.012 | 1.188 | 1.212 | #DIV/0!
|  |  | 42 | oC | 1.108 | kΩ | 1.097 | kΩ | 1.119 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 1.108 | 0.011 | 1.097 | 1.119 | #DIV/0!
| Pin 7 - Pin 8 |  | 30 | oC | 4.836 | kΩ | 4.785 | kΩ | 4.887 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 4.836 | 0.051 | 4.785 | 4.887 | #DIV/0!
| (Temp 700 T1) |  | 35 | oC | 3.918 | kΩ | 3.879 | kΩ | 3.957 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 3.918 | 0.039 | 3.879 | 3.957 | #DIV/0!
|  |  | 37 | oC | 3.61 | kΩ | 3.574 | kΩ | 3.646 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 3.61 | 0.036 | 3.574 | 3.646 | #DIV/0!
|  |  | 40 | oC | 3.196 | kΩ | 3.164 | kΩ | 3.228 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 3.196 | 0.032 | 3.164 | 3.228 | #DIV/0!
|  |  | 42 | oC | 2.95 | kΩ | 2.921 | kΩ | 2.979 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 2.95 | 0.029 | 2.921 | 2.979 | #DIV/0!
| Pin 7 - Pin 5 |  | 30 | oC | 24.27 | kΩ | 24.022 | kΩ | 24.518 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 24.27 | 0.248 | 24.022 | 24.518 | #DIV/0!
| (Temp 700 T2) |  | 35 | oC | 19.74 | kΩ | 19.545 | kΩ | 19.935 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 19.74 | 0.195 | 19.545 | 19.935 | #DIV/0!
|  |  | 37 | oC | 18.21 | kΩ | 18.03 | kΩ | 18.39 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 18.21 | 0.18 | 18.03 | 18.39 | #DIV/0!
|  |  | 40 | oC | 16.15 | kΩ | 15.992 | kΩ | 16.308 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 16.15 | 0.158 | 15.992 | 16.308 | #DIV/0!
|  |  | 42 | oC | 14.92 | kΩ | 14.777 | kΩ | 15.063 | kΩ |  |  |  | #DIV/0! | kΩ | #DIV/0! |  |  |  | 14.92 | 0.143 | 14.777 | 15.063 | #DIV/0!
6. Blood Pressure Sensitivity Test
| BP - 1 |  | 5 | μV/V | 10 | mV | 9.875 | mV | 10.125 | mV |  |  |  | #DIV/0! | mV | #DIV/0! |  |  |  | 10 | 0.125 | 9.875 | 10.125 | #DIV/0!
| (400 mmHg) |  | 40 | μV/V | 80 | mV | 79 | mV | 81 | mV |  |  |  | #DIV/0! | mV | #DIV/0! |  |  |  | 80 | 1 | 79 | 81 | #DIV/0!
| BP - 2 |  | 5 | μV/V | 6 | mV | 5.925 | mV | 6.075 | mV |  |  |  | #DIV/0! | mV | #DIV/0! |  |  |  | 6 | 0.075 | 5.925 | 6.075 | #DIV/0!
| (240 mmHg) |  | 40 | μV/V | 48 | mV | 47.42 | mV | 48.58 | mV |  |  |  | #DIV/0! | mV | #DIV/0! |  |  |  | 48 | 0.58 | 47.42 | 48.58 | #DIV/0!
Calibrated by : | [7 empty] | Checked by :
Date : |  | 2016-01-07 | [5 empty] | Date :